## Thermo Starting Settings
DIA and Chromatogram Library Settings for Thermo Instruments
For Searle et al, 2018 (https://www.nature.com/articles/s41467-018-07454-w)
Contact Brian Searle at searleb@uw.edu or bsearle@systemsbiology.org
These are meant to be starting settings and windowing schemes. You should optimize these settings for your system for best results. "24 mzol" shorthand means 24 m/z wide tiles that are staggered (overlapped) every other cycle. These settings assume that your peak widths are around 25-45 seconds wide. If your peaks are significantly outside that range then you should adjust the windowing scheme accordingly. If your peaks are wider then you can use narrower windows (e.g. if it says 24 mzol you might get better results with 16 mzol). If your peaks are narrower then you should use wider windows (e.g. if it says 8 mzol you might get better results with 12 mzol).
Single Injection Quantitative Acquisitions:
| QE | QE+ | QE-HF | QE-HFX | Fusion | Lumos
Windowing Scheme | 400-1000mz 24 mzol | 400-1000mz 24 mzol | 400-1000mz 24 mzol | 400-1000mz 8 mzol | 400-1000mz 24 mzol | 400-1000mz 8 mzol
Resolution | 30000 | 30000 | 30000 | 15000 | 30000 | 15000
Max IIT | 60 | 60 | 60 | 20 | 60 | 20
AGC Target | 1000000 | 1000000 | 1000000 | 1000000 | 400000 | 400000
Default Charge | 3 | 3 | 3 | 3 | 3 | 3
Loop Count | 25 | 25 | 25 | 75 | 25 | 75
(N)CE | 27 | 27 | 27 | 27 | 30 | 30
6x Injection Chromatogram Library Acquisitions:
| QE | QE+ | QE-HF | QE-HFX | Fusion | Lumos
Windowing Scheme | 6x 4 mzol | 6x 4 mzol | 6x 4 mzol | 6x 4 mzol | 6x 4 mzol | 6x 4 mzol
Resolution | 30000 | 30000 | 30000 | 30000 | 30000 | 30000
Max IIT | 60 | 60 | 60 | 60 | 60 | 60
AGC Target | 1000000 | 1000000 | 1000000 | 1000000 | 400000 | 400000
Default Charge | 3 | 3 | 3 | 3 | 3 | 3
Loop Count | 25 | 25 | 25 | 25 | 25 | 25
(N)CE | 27 | 27 | 27 | 27 | 30 | 30

## 400-1000mz 8 mzol
Scan | Start m/z | Center m/z | Stop m/z | Delta m/z
1 | 400.4319 | 404.4337 | 408.4355 | 8.0037
2 | 408.4355 | 412.4374 | 416.4392 | 8.0036
3 | 416.4392 | 420.441 | 424.4428 | 8.0036
4 | 424.4428 | 428.4446 | 432.4464 | 8.0036
5 | 432.4464 | 436.4483 | 440.4501 | 8.0037
6 | 440.4501 | 444.4519 | 448.4537 | 8.0036
7 | 448.4537 | 452.4556 | 456.4574 | 8.0036
8 | 456.4574 | 460.4592 | 464.461 | 8.0036
9 | 464.461 | 468.4628 | 472.4646 | 8.0036
10 | 472.4646 | 476.4665 | 480.4683 | 8.0037
11 | 480.4683 | 484.4701 | 488.4719 | 8.0036
12 | 488.4719 | 492.4737 | 496.4756 | 8.0036
13 | 496.4756 | 500.4774 | 504.4792 | 8.0036
14 | 504.4792 | 508.481 | 512.4828 | 8.0037
15 | 512.4828 | 516.4846 | 520.4864 | 8.0036
16 | 520.4864 | 524.4883 | 528.4901 | 8.0037
17 | 528.4901 | 532.4919 | 536.4938 | 8.0037
18 | 536.4938 | 540.4956 | 544.4974 | 8.0036
19 | 544.4974 | 548.4992 | 552.501 | 8.0037
20 | 552.501 | 556.5028 | 560.5046 | 8.0036
21 | 560.5046 | 564.5065 | 568.5083 | 8.0037
22 | 568.5083 | 572.5101 | 576.512 | 8.0037
23 | 576.512 | 580.5138 | 584.5156 | 8.0036
24 | 584.5156 | 588.5174 | 592.5192 | 8.0037
25 | 592.5192 | 596.521 | 600.5228 | 8.0036
26 | 600.5228 | 604.5247 | 608.5265 | 8.0037
27 | 608.5265 | 612.5283 | 616.5302 | 8.0037
28 | 616.5302 | 620.532 | 624.5338 | 8.0036
29 | 624.5338 | 628.5356 | 632.5374 | 8.0037
30 | 632.5374 | 636.5392 | 640.541 | 8.0036
31 | 640.541 | 644.5429 | 648.5447 | 8.0037
32 | 648.5447 | 652.5465 | 656.5483 | 8.0037
33 | 656.5483 | 660.5502 | 664.5519 | 8.0036
34 | 664.5519 | 668.5538 | 672.5556 | 8.0037
35 | 672.5556 | 676.5574 | 680.5592 | 8.0036
36 | 680.5592 | 684.561 | 688.5629 | 8.0037
37 | 688.5629 | 692.5647 | 696.5665 | 8.0037
38 | 696.5665 | 700.5684 | 704.5701 | 8.0036
39 | 704.5701 | 708.572 | 712.5738 | 8.0037
40 | 712.5738 | 716.5756 | 720.5774 | 8.0036
41 | 720.5774 | 724.5792 | 728.5811 | 8.0037
42 | 728.5811 | 732.5829 | 736.5847 | 8.0037
43 | 736.5847 | 740.5865 | 744.5883 | 8.0036
44 | 744.5883 | 748.5901 | 752.592 | 8.0037
45 | 752.592 | 756.5938 | 760.5956 | 8.0036
46 | 760.5956 | 764.5974 | 768.5992 | 8.0037
47 | 768.5992 | 772.6011 | 776.6029 | 8.0037
48 | 776.6029 | 780.6047 | 784.6065 | 8.0036
49 | 784.6065 | 788.6083 | 792.6102 | 8.0037
50 | 792.6102 | 796.6119 | 800.6138 | 8.0036
51 | 800.6138 | 804.6156 | 808.6174 | 8.0037
52 | 808.6174 | 812.6193 | 816.6211 | 8.0037
53 | 816.6211 | 820.6229 | 824.6247 | 8.0036
54 | 824.6247 | 828.6265 | 832.6284 | 8.0037
55 | 832.6284 | 836.6302 | 840.632 | 8.0037
56 | 840.632 | 844.6338 | 848.6356 | 8.0036
57 | 848.6356 | 852.6374 | 856.6393 | 8.0037
58 | 856.6393 | 860.6411 | 864.6429 | 8.0036
59 | 864.6429 | 868.6447 | 872.6466 | 8.0037
60 | 872.6466 | 876.6484 | 880.6502 | 8.0037
61 | 880.6502 | 884.652 | 888.6538 | 8.0036
62 | 888.6538 | 892.6556 | 896.6575 | 8.0037
63 | 896.6575 | 900.6593 | 904.6611 | 8.0036
64 | 904.6611 | 908.6629 | 912.6647 | 8.0037
65 | 912.6647 | 916.6666 | 920.6684 | 8.0037
66 | 920.6684 | 924.6702 | 928.672 | 8.0036
67 | 928.672 | 932.6738 | 936.6757 | 8.0037
68 | 936.6757 | 940.6775 | 944.6793 | 8.0036
69 | 944.6793 | 948.6811 | 952.6829 | 8.0037
70 | 952.6829 | 956.6848 | 960.6866 | 8.0037
71 | 960.6866 | 964.6884 | 968.6902 | 8.0036
72 | 968.6902 | 972.692 | 976.6938 | 8.0037
73 | 976.6938 | 980.6957 | 984.6975 | 8.0036
74 | 984.6975 | 988.6993 | 992.7011 | 8.0037
75 | 992.7011 | 996.7029 | 1000.7048 | 8.0037
76 | 396.4301 | 400.4319 | 404.4337 | 8.0036
77 | 404.4337 | 408.4355 | 412.4373 | 8.0036
78 | 412.4373 | 416.4392 | 420.441 | 8.0037
79 | 420.441 | 424.4428 | 428.4446 | 8.0036
80 | 428.4446 | 432.4465 | 436.4483 | 8.0036
81 | 436.4483 | 440.4501 | 444.4519 | 8.0036
82 | 444.4519 | 448.4537 | 452.4555 | 8.0036
83 | 452.4555 | 456.4574 | 460.4592 | 8.0037
84 | 460.4592 | 464.461 | 468.4628 | 8.0036
85 | 468.4628 | 472.4647 | 476.4665 | 8.0036
86 | 476.4665 | 480.4683 | 484.4701 | 8.0036
87 | 484.4701 | 488.4719 | 492.4737 | 8.0036
88 | 492.4737 | 496.4756 | 500.4774 | 8.0037
89 | 500.4774 | 504.4792 | 508.481 | 8.0036
90 | 508.481 | 512.4828 | 516.4847 | 8.0037
91 | 516.4847 | 520.4864 | 524.4883 | 8.0036
92 | 524.4883 | 528.4901 | 532.4919 | 8.0037
93 | 532.4919 | 536.4938 | 540.4955 | 8.0036
94 | 540.4955 | 544.4974 | 548.4992 | 8.0037
95 | 548.4992 | 552.501 | 556.5029 | 8.0037
96 | 556.5029 | 560.5046 | 564.5065 | 8.0036
97 | 564.5065 | 568.5083 | 572.5101 | 8.0037
98 | 572.5101 | 576.512 | 580.5137 | 8.0036
99 | 580.5137 | 584.5156 | 588.5174 | 8.0037
100 | 588.5174 | 592.5192 | 596.5211 | 8.0037
101 | 596.5211 | 600.5228 | 604.5247 | 8.0036
102 | 604.5247 | 608.5265 | 612.5283 | 8.0037
103 | 612.5283 | 616.5302 | 620.5319 | 8.0036
104 | 620.5319 | 624.5338 | 628.5356 | 8.0037
105 | 628.5356 | 632.5374 | 636.5393 | 8.0037
106 | 636.5393 | 640.541 | 644.5429 | 8.0036
107 | 644.5429 | 648.5447 | 652.5465 | 8.0037
108 | 652.5465 | 656.5483 | 660.5501 | 8.0036
109 | 660.5501 | 664.5519 | 668.5538 | 8.0037
110 | 668.5538 | 672.5556 | 676.5574 | 8.0037
111 | 676.5574 | 680.5592 | 684.561 | 8.0036
112 | 684.561 | 688.5629 | 692.5647 | 8.0037
113 | 692.5647 | 696.5665 | 700.5683 | 8.0036
114 | 700.5683 | 704.5701 | 708.572 | 8.0037
115 | 708.572 | 712.5738 | 716.5756 | 8.0037
116 | 716.5756 | 720.5774 | 724.5792 | 8.0036
117 | 724.5792 | 728.5811 | 732.5829 | 8.0037
118 | 732.5829 | 736.5847 | 740.5865 | 8.0036
119 | 740.5865 | 744.5883 | 748.5901 | 8.0037
120 | 748.5901 | 752.592 | 756.5938 | 8.0037
121 | 756.5938 | 760.5956 | 764.5974 | 8.0036
122 | 764.5974 | 768.5992 | 772.6011 | 8.0037
123 | 772.6011 | 776.6029 | 780.6047 | 8.0036
124 | 780.6047 | 784.6065 | 788.6083 | 8.0037
125 | 788.6083 | 792.6102 | 796.612 | 8.0037
126 | 796.612 | 800.6138 | 804.6156 | 8.0036
127 | 804.6156 | 808.6174 | 812.6193 | 8.0037
128 | 812.6193 | 816.6211 | 820.6229 | 8.0037
129 | 820.6229 | 824.6248 | 828.6265 | 8.0036
130 | 828.6265 | 832.6284 | 836.6302 | 8.0037
131 | 836.6302 | 840.632 | 844.6338 | 8.0036
132 | 844.6338 | 848.6356 | 852.6374 | 8.0037
133 | 852.6374 | 856.6393 | 860.6411 | 8.0037
134 | 860.6411 | 864.6429 | 868.6447 | 8.0036
135 | 868.6447 | 872.6466 | 876.6484 | 8.0037
136 | 876.6484 | 880.6502 | 884.652 | 8.0036
137 | 884.652 | 888.6538 | 892.6556 | 8.0037
138 | 892.6556 | 896.6575 | 900.6593 | 8.0037
139 | 900.6593 | 904.6611 | 908.6629 | 8.0036
140 | 908.6629 | 912.6647 | 916.6666 | 8.0037
141 | 916.6666 | 920.6683 | 924.6702 | 8.0036
142 | 924.6702 | 928.672 | 932.6738 | 8.0037
143 | 932.6738 | 936.6757 | 940.6775 | 8.0037
144 | 940.6775 | 944.6793 | 948.6811 | 8.0036
145 | 948.6811 | 952.6829 | 956.6848 | 8.0037
146 | 956.6848 | 960.6865 | 964.6884 | 8.0036
147 | 964.6884 | 968.6902 | 972.692 | 8.0037
148 | 972.692 | 976.6938 | 980.6957 | 8.0037
149 | 980.6957 | 984.6975 | 988.6993 | 8.0036
150 | 988.6993 | 992.7011 | 996.7029 | 8.0037
151 | 996.7029 | 1000.7047 | 1004.7065 | 8.0036

## 400-1000mz 12 mzol
Scan | Start m/z | Center m/z | Stop m/z | Delta m/z
1 | 400.4319 | 406.4346 | 412.4373 | 12.0055
2 | 412.4373 | 418.4401 | 424.4428 | 12.0055
3 | 424.4428 | 430.4456 | 436.4483 | 12.0055
4 | 436.4483 | 442.451 | 448.4537 | 12.0055
5 | 448.4537 | 454.4565 | 460.4592 | 12.0055
6 | 460.4592 | 466.4619 | 472.4646 | 12.0054
7 | 472.4646 | 478.4674 | 484.4701 | 12.0055
8 | 484.4701 | 490.4728 | 496.4756 | 12.0055
9 | 496.4756 | 502.4783 | 508.481 | 12.0055
10 | 508.481 | 514.4838 | 520.4864 | 12.0054
11 | 520.4864 | 526.4892 | 532.4919 | 12.0055
12 | 532.4919 | 538.4946 | 544.4974 | 12.0054
13 | 544.4974 | 550.5001 | 556.5029 | 12.0055
14 | 556.5029 | 562.5056 | 568.5083 | 12.0054
15 | 568.5083 | 574.511 | 580.5137 | 12.0054
16 | 580.5137 | 586.5165 | 592.5192 | 12.0055
17 | 592.5192 | 598.522 | 604.5247 | 12.0054
18 | 604.5247 | 610.5274 | 616.5302 | 12.0055
19 | 616.5302 | 622.5328 | 628.5356 | 12.0054
20 | 628.5356 | 634.5383 | 640.541 | 12.0054
21 | 640.541 | 646.5438 | 652.5465 | 12.0055
22 | 652.5465 | 658.5492 | 664.5519 | 12.0054
23 | 664.5519 | 670.5547 | 676.5574 | 12.0055
24 | 676.5574 | 682.5602 | 688.5629 | 12.0054
25 | 688.5629 | 694.5656 | 700.5683 | 12.0054
26 | 700.5683 | 706.571 | 712.5738 | 12.0055
27 | 712.5738 | 718.5765 | 724.5792 | 12.0054
28 | 724.5792 | 730.582 | 736.5847 | 12.0055
29 | 736.5847 | 742.5874 | 748.5901 | 12.0054
30 | 748.5901 | 754.5929 | 760.5956 | 12.0054
31 | 760.5956 | 766.5983 | 772.6011 | 12.0055
32 | 772.6011 | 778.6038 | 784.6065 | 12.0054
33 | 784.6065 | 790.6092 | 796.612 | 12.0055
34 | 796.612 | 802.6147 | 808.6174 | 12.0054
35 | 808.6174 | 814.6202 | 820.6229 | 12.0055
36 | 820.6229 | 826.6256 | 832.6284 | 12.0054
37 | 832.6284 | 838.6311 | 844.6338 | 12.0054
38 | 844.6338 | 850.6365 | 856.6393 | 12.0055
39 | 856.6393 | 862.642 | 868.6447 | 12.0054
40 | 868.6447 | 874.6475 | 880.6502 | 12.0055
41 | 880.6502 | 886.653 | 892.6556 | 12.0054
42 | 892.6556 | 898.6583 | 904.6611 | 12.0054
43 | 904.6611 | 910.6638 | 916.6666 | 12.0055
44 | 916.6666 | 922.6693 | 928.672 | 12.0054
45 | 928.672 | 934.6747 | 940.6775 | 12.0055
46 | 940.6775 | 946.6802 | 952.6829 | 12.0054
47 | 952.6829 | 958.6857 | 964.6884 | 12.0054
48 | 964.6884 | 970.6911 | 976.6938 | 12.0055
49 | 976.6938 | 982.6965 | 988.6993 | 12.0054
50 | 988.6993 | 994.702 | 1000.7048 | 12.0055
51 | 394.4292 | 400.4319 | 406.4346 | 12.0055
52 | 406.4346 | 412.4374 | 418.4401 | 12.0055
53 | 418.4401 | 424.4428 | 430.4456 | 12.0055
54 | 430.4456 | 436.4483 | 442.451 | 12.0054
55 | 442.451 | 448.4537 | 454.4565 | 12.0055
56 | 454.4565 | 460.4592 | 466.4619 | 12.0055
57 | 466.4619 | 472.4647 | 478.4674 | 12.0055
58 | 478.4674 | 484.4701 | 490.4728 | 12.0055
59 | 490.4728 | 496.4756 | 502.4783 | 12.0054
60 | 502.4783 | 508.481 | 514.4838 | 12.0055
61 | 514.4838 | 520.4864 | 526.4892 | 12.0054
62 | 526.4892 | 532.4919 | 538.4946 | 12.0054
63 | 538.4946 | 544.4974 | 550.5001 | 12.0055
64 | 550.5001 | 556.5028 | 562.5055 | 12.0054
65 | 562.5055 | 568.5083 | 574.511 | 12.0055
66 | 574.511 | 580.5138 | 586.5165 | 12.0054
67 | 586.5165 | 592.5192 | 598.5219 | 12.0054
68 | 598.5219 | 604.5247 | 610.5274 | 12.0055
69 | 610.5274 | 616.5302 | 622.5328 | 12.0054
70 | 622.5328 | 628.5356 | 634.5383 | 12.0055
71 | 634.5383 | 640.541 | 646.5438 | 12.0054
72 | 646.5438 | 652.5465 | 658.5493 | 12.0055
73 | 658.5493 | 664.552 | 670.5547 | 12.0054
74 | 670.5547 | 676.5574 | 682.5601 | 12.0054
75 | 682.5601 | 688.5629 | 694.5656 | 12.0055
76 | 694.5656 | 700.5684 | 706.571 | 12.0054
77 | 706.571 | 712.5738 | 718.5765 | 12.0055
78 | 718.5765 | 724.5792 | 730.582 | 12.0054
79 | 730.582 | 736.5847 | 742.5874 | 12.0054
80 | 742.5874 | 748.5901 | 754.5929 | 12.0055
81 | 754.5929 | 760.5956 | 766.5983 | 12.0054
82 | 766.5983 | 772.6011 | 778.6038 | 12.0055
83 | 778.6038 | 784.6066 | 790.6092 | 12.0054
84 | 790.6092 | 796.6119 | 802.6147 | 12.0054
85 | 802.6147 | 808.6174 | 814.6202 | 12.0055
86 | 814.6202 | 820.6229 | 826.6256 | 12.0054
87 | 826.6256 | 832.6284 | 838.6311 | 12.0055
88 | 838.6311 | 844.6338 | 850.6365 | 12.0054
89 | 850.6365 | 856.6393 | 862.642 | 12.0054
90 | 862.642 | 868.6447 | 874.6475 | 12.0055
91 | 874.6475 | 880.6502 | 886.6529 | 12.0054
92 | 886.6529 | 892.6556 | 898.6584 | 12.0055
93 | 898.6584 | 904.6611 | 910.6638 | 12.0054
94 | 910.6638 | 916.6665 | 922.6693 | 12.0054
95 | 922.6693 | 928.672 | 934.6747 | 12.0055
96 | 934.6747 | 940.6775 | 946.6802 | 12.0054
97 | 946.6802 | 952.6829 | 958.6857 | 12.0055
98 | 958.6857 | 964.6884 | 970.6911 | 12.0054
99 | 970.6911 | 976.6938 | 982.6966 | 12.0055
100 | 982.6966 | 988.6993 | 994.702 | 12.0054
101 | 994.702 | 1000.7047 | 1006.7075 | 12.0054

## 400-1000mz 16 mzol
Scan | Start m/z | Center m/z | Stop m/z | Delta m/z
1 | 400.4319 | 408.4355 | 416.4392 | 16.0073
2 | 416.4392 | 424.4428 | 432.4464 | 16.0073
3 | 432.4464 | 440.4501 | 448.4537 | 16.0073
4 | 448.4537 | 456.4574 | 464.461 | 16.0073
5 | 464.461 | 472.4647 | 480.4683 | 16.0073
6 | 480.4683 | 488.4719 | 496.4756 | 16.0073
7 | 496.4756 | 504.4792 | 512.4828 | 16.0073
8 | 512.4828 | 520.4864 | 528.4901 | 16.0073
9 | 528.4901 | 536.4938 | 544.4974 | 16.0073
10 | 544.4974 | 552.501 | 560.5046 | 16.0073
11 | 560.5046 | 568.5083 | 576.512 | 16.0073
12 | 576.512 | 584.5156 | 592.5192 | 16.0073
13 | 592.5192 | 600.5228 | 608.5265 | 16.0073
14 | 608.5265 | 616.5302 | 624.5338 | 16.0073
15 | 624.5338 | 632.5373 | 640.541 | 16.0073
16 | 640.541 | 648.5447 | 656.5483 | 16.0073
17 | 656.5483 | 664.552 | 672.5556 | 16.0073
18 | 672.5556 | 680.5592 | 688.5629 | 16.0073
19 | 688.5629 | 696.5665 | 704.5701 | 16.0073
20 | 704.5701 | 712.5737 | 720.5774 | 16.0073
21 | 720.5774 | 728.5811 | 736.5847 | 16.0073
22 | 736.5847 | 744.5884 | 752.592 | 16.0073
23 | 752.592 | 760.5956 | 768.5992 | 16.0073
24 | 768.5992 | 776.6029 | 784.6065 | 16.0073
25 | 784.6065 | 792.6101 | 800.6138 | 16.0073
26 | 800.6138 | 808.6174 | 816.6211 | 16.0073
27 | 816.6211 | 824.6248 | 832.6284 | 16.0073
28 | 832.6284 | 840.632 | 848.6356 | 16.0073
29 | 848.6356 | 856.6393 | 864.6429 | 16.0073
30 | 864.6429 | 872.6466 | 880.6502 | 16.0073
31 | 880.6502 | 888.6538 | 896.6575 | 16.0073
32 | 896.6575 | 904.6611 | 912.6647 | 16.0073
33 | 912.6647 | 920.6683 | 928.672 | 16.0073
34 | 928.672 | 936.6757 | 944.6793 | 16.0073
35 | 944.6793 | 952.6829 | 960.6866 | 16.0073
36 | 960.6866 | 968.6902 | 976.6938 | 16.0073
37 | 976.6938 | 984.6975 | 992.7011 | 16.0073
38 | 992.7011 | 1000.7047 | 1008.7084 | 16.0073
39 | 392.4282 | 400.4319 | 408.4355 | 16.0073
40 | 408.4355 | 416.4392 | 424.4428 | 16.0073
41 | 424.4428 | 432.4465 | 440.4501 | 16.0073
42 | 440.4501 | 448.4537 | 456.4574 | 16.0073
43 | 456.4574 | 464.461 | 472.4646 | 16.0073
44 | 472.4646 | 480.4683 | 488.4719 | 16.0073
45 | 488.4719 | 496.4756 | 504.4792 | 16.0073
46 | 504.4792 | 512.4828 | 520.4864 | 16.0073
47 | 520.4864 | 528.4901 | 536.4938 | 16.0073
48 | 536.4938 | 544.4974 | 552.501 | 16.0073
49 | 552.501 | 560.5046 | 568.5083 | 16.0073
50 | 568.5083 | 576.512 | 584.5156 | 16.0073
51 | 584.5156 | 592.5192 | 600.5228 | 16.0073
52 | 600.5228 | 608.5265 | 616.5302 | 16.0073
53 | 616.5302 | 624.5338 | 632.5374 | 16.0073
54 | 632.5374 | 640.541 | 648.5447 | 16.0073
55 | 648.5447 | 656.5483 | 664.5519 | 16.0073
56 | 664.5519 | 672.5555 | 680.5592 | 16.0073
57 | 680.5592 | 688.5629 | 696.5665 | 16.0073
58 | 696.5665 | 704.5702 | 712.5738 | 16.0073
59 | 712.5738 | 720.5774 | 728.5811 | 16.0073
60 | 728.5811 | 736.5847 | 744.5883 | 16.0073
61 | 744.5883 | 752.5919 | 760.5956 | 16.0073
62 | 760.5956 | 768.5992 | 776.6029 | 16.0073
63 | 776.6029 | 784.6066 | 792.6102 | 16.0073
64 | 792.6102 | 800.6138 | 808.6174 | 16.0073
65 | 808.6174 | 816.6211 | 824.6247 | 16.0073
66 | 824.6247 | 832.6284 | 840.632 | 16.0073
67 | 840.632 | 848.6356 | 856.6393 | 16.0073
68 | 856.6393 | 864.6429 | 872.6466 | 16.0073
69 | 872.6466 | 880.6502 | 888.6538 | 16.0073
70 | 888.6538 | 896.6575 | 904.6611 | 16.0073
71 | 904.6611 | 912.6647 | 920.6684 | 16.0073
72 | 920.6684 | 928.672 | 936.6757 | 16.0073
73 | 936.6757 | 944.6793 | 952.6829 | 16.0073
74 | 952.6829 | 960.6865 | 968.6902 | 16.0073
75 | 968.6902 | 976.6938 | 984.6975 | 16.0073
76 | 984.6975 | 992.7011 | 1000.7048 | 16.0073

## 400-1000mz 24 mzol
Scan | Start m/z | Center m/z | Stop m/z | Delta m/z
1 | 400.4319 | 412.4373 | 424.4428 | 24.0109
2 | 424.4428 | 436.4483 | 448.4537 | 24.0109
3 | 448.4537 | 460.4592 | 472.4646 | 24.0109
4 | 472.4646 | 484.4701 | 496.4756 | 24.0109
5 | 496.4756 | 508.481 | 520.4864 | 24.0109
6 | 520.4864 | 532.4919 | 544.4974 | 24.0109
7 | 544.4974 | 556.5028 | 568.5083 | 24.0109
8 | 568.5083 | 580.5138 | 592.5192 | 24.0109
9 | 592.5192 | 604.5247 | 616.5302 | 24.0109
10 | 616.5302 | 628.5356 | 640.541 | 24.0109
11 | 640.541 | 652.5465 | 664.5519 | 24.0109
12 | 664.5519 | 676.5574 | 688.5629 | 24.0109
13 | 688.5629 | 700.5684 | 712.5738 | 24.0109
14 | 712.5738 | 724.5792 | 736.5847 | 24.0109
15 | 736.5847 | 748.5901 | 760.5956 | 24.0109
16 | 760.5956 | 772.6011 | 784.6065 | 24.0109
17 | 784.6065 | 796.6119 | 808.6174 | 24.0109
18 | 808.6174 | 820.6229 | 832.6284 | 24.0109
19 | 832.6284 | 844.6338 | 856.6393 | 24.0109
20 | 856.6393 | 868.6448 | 880.6502 | 24.0109
21 | 880.6502 | 892.6556 | 904.6611 | 24.0109
22 | 904.6611 | 916.6665 | 928.672 | 24.0109
23 | 928.672 | 940.6775 | 952.6829 | 24.0109
24 | 952.6829 | 964.6884 | 976.6938 | 24.0109
25 | 976.6938 | 988.6993 | 1000.7048 | 24.0109
26 | 388.4264 | 400.4319 | 412.4373 | 24.0109
27 | 412.4373 | 424.4428 | 436.4483 | 24.0109
28 | 436.4483 | 448.4537 | 460.4592 | 24.0109
29 | 460.4592 | 472.4647 | 484.4701 | 24.0109
30 | 484.4701 | 496.4756 | 508.481 | 24.0109
31 | 508.481 | 520.4864 | 532.4919 | 24.0109
32 | 532.4919 | 544.4974 | 556.5029 | 24.0109
33 | 556.5029 | 568.5083 | 580.5137 | 24.0109
34 | 580.5137 | 592.5192 | 604.5247 | 24.0109
35 | 604.5247 | 616.5302 | 628.5356 | 24.0109
36 | 628.5356 | 640.541 | 652.5465 | 24.0109
37 | 652.5465 | 664.552 | 676.5574 | 24.0109
38 | 676.5574 | 688.5629 | 700.5683 | 24.0109
39 | 700.5683 | 712.5737 | 724.5792 | 24.0109
40 | 724.5792 | 736.5847 | 748.5901 | 24.0109
41 | 748.5901 | 760.5956 | 772.6011 | 24.0109
42 | 772.6011 | 784.6066 | 796.612 | 24.0109
43 | 796.612 | 808.6174 | 820.6229 | 24.0109
44 | 820.6229 | 832.6284 | 844.6338 | 24.0109
45 | 844.6338 | 856.6393 | 868.6447 | 24.0109
46 | 868.6447 | 880.6502 | 892.6556 | 24.0109
47 | 892.6556 | 904.6611 | 916.6666 | 24.0109
48 | 916.6666 | 928.672 | 940.6775 | 24.0109
49 | 940.6775 | 952.6829 | 964.6884 | 24.0109
50 | 964.6884 | 976.6938 | 988.6993 | 24.0109
51 | 988.6993 | 1000.7047 | 1012.7102 | 24.0109

## Chromatogram Library Scheme
400-500 m/z chromatogram library acquisition |  |  |  |  | 500-600 m/z chromatogram library acquisition |  |  |  |  | 600-700 m/z chromatogram library acquisition |  |  |  |  | 700-800 m/z chromatogram library acquisition |  |  |  |  | 800-900 m/z chromatogram library acquisition |  |  |  |  | 900-1000 m/z chromatogram library acquisition
Scan | Start m/z | Center m/z | Stop m/z | Delta m/z | Scan | Start m/z | Center m/z | Stop m/z | Delta m/z | Scan | Start m/z | Center m/z | Stop m/z | Delta m/z | Scan | Start m/z | Center m/z | Stop m/z | Delta m/z | Scan | Start m/z | Center m/z | Stop m/z | Delta m/z | Scan | Start m/z | Center m/z | Stop m/z | Delta m/z
1 | 400.4319 | 402.4328 | 404.4337 | 4.0018 | 1 | 500.4774 | 502.4783 | 504.4792 | 4.0018 | 1 | 600.5228 | 602.5237 | 604.5247 | 4.0018 | 1 | 700.5683 | 702.5692 | 704.5701 | 4.0018 | 1 | 800.6138 | 802.6147 | 804.6156 | 4.0018 | 1 | 900.6593 | 902.6602 | 904.6611 | 4.0018
2 | 404.4337 | 406.4346 | 408.4355 | 4.0018 | 2 | 504.4792 | 506.4801 | 508.481 | 4.0018 | 2 | 604.5247 | 606.5256 | 608.5265 | 4.0018 | 2 | 704.5701 | 706.571 | 708.572 | 4.0018 | 2 | 804.6156 | 806.6165 | 808.6174 | 4.0018 | 2 | 904.6611 | 906.662 | 908.6629 | 4.0018
3 | 408.4355 | 410.4365 | 412.4373 | 4.0018 | 3 | 508.481 | 510.4819 | 512.4828 | 4.0018 | 3 | 608.5265 | 610.5274 | 612.5283 | 4.0018 | 3 | 708.572 | 710.5729 | 712.5738 | 4.0018 | 3 | 808.6174 | 810.6183 | 812.6193 | 4.0018 | 3 | 908.6629 | 910.6638 | 912.6647 | 4.0018
4 | 412.4373 | 414.4383 | 416.4392 | 4.0018 | 4 | 512.4828 | 514.4838 | 516.4847 | 4.0018 | 4 | 612.5283 | 614.5292 | 616.5302 | 4.0018 | 4 | 712.5738 | 714.5747 | 716.5756 | 4.0018 | 4 | 812.6193 | 814.6202 | 816.6211 | 4.0018 | 4 | 912.6647 | 914.6657 | 916.6666 | 4.0018
5 | 416.4392 | 418.4401 | 420.441 | 4.0018 | 5 | 516.4847 | 518.4856 | 520.4864 | 4.0018 | 5 | 616.5302 | 618.531 | 620.5319 | 4.0018 | 5 | 716.5756 | 718.5765 | 720.5774 | 4.0018 | 5 | 816.6211 | 818.622 | 820.6229 | 4.0018 | 5 | 916.6666 | 918.6675 | 920.6684 | 4.0018
6 | 420.441 | 422.4419 | 424.4428 | 4.0018 | 6 | 520.4864 | 522.4874 | 524.4883 | 4.0018 | 6 | 620.5319 | 622.5328 | 624.5338 | 4.0018 | 6 | 720.5774 | 722.5783 | 724.5792 | 4.0018 | 6 | 820.6229 | 822.6238 | 824.6247 | 4.0018 | 6 | 920.6684 | 922.6693 | 924.6702 | 4.0018
7 | 424.4428 | 426.4437 | 428.4446 | 4.0018 | 7 | 524.4883 | 526.4892 | 528.4901 | 4.0018 | 7 | 624.5338 | 626.5347 | 628.5356 | 4.0018 | 7 | 724.5792 | 726.5801 | 728.5811 | 4.0018 | 7 | 824.6247 | 826.6256 | 828.6265 | 4.0018 | 7 | 924.6702 | 926.6711 | 928.672 | 4.0018
8 | 428.4446 | 430.4456 | 432.4464 | 4.0018 | 8 | 528.4901 | 530.491 | 532.4919 | 4.0018 | 8 | 628.5356 | 630.5365 | 632.5374 | 4.0018 | 8 | 728.5811 | 730.582 | 732.5829 | 4.0018 | 8 | 828.6265 | 830.6274 | 832.6284 | 4.0018 | 8 | 928.672 | 930.6729 | 932.6738 | 4.0018
9 | 432.4464 | 434.4474 | 436.4483 | 4.0018 | 9 | 532.4919 | 534.4929 | 536.4938 | 4.0018 | 9 | 632.5374 | 634.5383 | 636.5393 | 4.0018 | 9 | 732.5829 | 734.5838 | 736.5847 | 4.0018 | 9 | 832.6284 | 834.6293 | 836.6302 | 4.0018 | 9 | 932.6738 | 934.6747 | 936.6757 | 4.0018
10 | 436.4483 | 438.4492 | 440.4501 | 4.0018 | 10 | 536.4938 | 538.4946 | 540.4955 | 4.0018 | 10 | 636.5393 | 638.5402 | 640.541 | 4.0018 | 10 | 736.5847 | 738.5856 | 740.5865 | 4.0018 | 10 | 836.6302 | 838.6311 | 840.632 | 4.0018 | 10 | 936.6757 | 938.6766 | 940.6775 | 4.0018
11 | 440.4501 | 442.451 | 444.4519 | 4.0018 | 11 | 540.4955 | 542.4965 | 544.4974 | 4.0018 | 11 | 640.541 | 642.5419 | 644.5429 | 4.0018 | 11 | 740.5865 | 742.5874 | 744.5883 | 4.0018 | 11 | 840.632 | 842.6329 | 844.6338 | 4.0018 | 11 | 940.6775 | 942.6783 | 944.6793 | 4.0018
12 | 444.4519 | 446.4528 | 448.4537 | 4.0018 | 12 | 544.4974 | 546.4983 | 548.4992 | 4.0018 | 12 | 644.5429 | 646.5438 | 648.5447 | 4.0018 | 12 | 744.5883 | 746.5892 | 748.5901 | 4.0018 | 12 | 844.6338 | 846.6347 | 848.6356 | 4.0018 | 12 | 944.6793 | 946.6802 | 948.6811 | 4.0018
13 | 448.4537 | 450.4547 | 452.4555 | 4.0018 | 13 | 548.4992 | 550.5001 | 552.501 | 4.0018 | 13 | 648.5447 | 650.5456 | 652.5465 | 4.0018 | 13 | 748.5901 | 750.5911 | 752.592 | 4.0018 | 13 | 848.6356 | 850.6365 | 852.6374 | 4.0018 | 13 | 948.6811 | 950.682 | 952.6829 | 4.0018
14 | 452.4555 | 454.4565 | 456.4574 | 4.0018 | 14 | 552.501 | 554.5019 | 556.5029 | 4.0018 | 14 | 652.5465 | 654.5474 | 656.5483 | 4.0018 | 14 | 752.592 | 754.5929 | 756.5938 | 4.0018 | 14 | 852.6374 | 854.6384 | 856.6393 | 4.0018 | 14 | 952.6829 | 954.6838 | 956.6848 | 4.0018
15 | 456.4574 | 458.4583 | 460.4592 | 4.0018 | 15 | 556.5029 | 558.5038 | 560.5046 | 4.0018 | 15 | 656.5483 | 658.5492 | 660.5501 | 4.0018 | 15 | 756.5938 | 758.5947 | 760.5956 | 4.0018 | 15 | 856.6393 | 858.6402 | 860.6411 | 4.0018 | 15 | 956.6848 | 958.6857 | 960.6866 | 4.0018
16 | 460.4592 | 462.4601 | 464.461 | 4.0018 | 16 | 560.5046 | 562.5055 | 564.5065 | 4.0018 | 16 | 660.5501 | 662.551 | 664.5519 | 4.0018 | 16 | 760.5956 | 762.5965 | 764.5974 | 4.0018 | 16 | 860.6411 | 862.642 | 864.6429 | 4.0018 | 16 | 960.6866 | 962.6875 | 964.6884 | 4.0018
17 | 464.461 | 466.4619 | 468.4628 | 4.0018 | 17 | 564.5065 | 566.5074 | 568.5083 | 4.0018 | 17 | 664.5519 | 666.5529 | 668.5538 | 4.0018 | 17 | 764.5974 | 766.5983 | 768.5992 | 4.0018 | 17 | 864.6429 | 866.6438 | 868.6447 | 4.0018 | 17 | 964.6884 | 966.6893 | 968.6902 | 4.0018
18 | 468.4628 | 470.4638 | 472.4646 | 4.0018 | 18 | 568.5083 | 570.5092 | 572.5101 | 4.0018 | 18 | 668.5538 | 670.5547 | 672.5556 | 4.0018 | 18 | 768.5992 | 770.6002 | 772.6011 | 4.0018 | 18 | 868.6447 | 870.6456 | 872.6466 | 4.0018 | 18 | 968.6902 | 970.6911 | 972.692 | 4.0018
19 | 472.4646 | 474.4656 | 476.4665 | 4.0018 | 19 | 572.5101 | 574.511 | 576.512 | 4.0018 | 19 | 672.5556 | 674.5565 | 676.5574 | 4.0018 | 19 | 772.6011 | 774.602 | 776.6029 | 4.0018 | 19 | 872.6466 | 874.6475 | 876.6484 | 4.0018 | 19 | 972.692 | 974.6929 | 976.6938 | 4.0018
20 | 476.4665 | 478.4674 | 480.4683 | 4.0018 | 20 | 576.512 | 578.5128 | 580.5137 | 4.0018 | 20 | 676.5574 | 678.5584 | 680.5592 | 4.0018 | 20 | 776.6029 | 778.6038 | 780.6047 | 4.0018 | 20 | 876.6484 | 878.6493 | 880.6502 | 4.0018 | 20 | 976.6938 | 978.6948 | 980.6957 | 4.0018
21 | 480.4683 | 482.4692 | 484.4701 | 4.0018 | 21 | 580.5137 | 582.5146 | 584.5156 | 4.0018 | 21 | 680.5592 | 682.5601 | 684.561 | 4.0018 | 21 | 780.6047 | 782.6056 | 784.6065 | 4.0018 | 21 | 880.6502 | 882.6511 | 884.652 | 4.0018 | 21 | 980.6957 | 982.6965 | 984.6975 | 4.0018
22 | 484.4701 | 486.471 | 488.4719 | 4.0018 | 22 | 584.5156 | 586.5165 | 588.5174 | 4.0018 | 22 | 684.561 | 686.562 | 688.5629 | 4.0018 | 22 | 784.6065 | 786.6074 | 788.6083 | 4.0018 | 22 | 884.652 | 886.6529 | 888.6538 | 4.0018 | 22 | 984.6975 | 986.6984 | 988.6993 | 4.0018
23 | 488.4719 | 490.4728 | 492.4737 | 4.0018 | 23 | 588.5174 | 590.5183 | 592.5192 | 4.0018 | 23 | 688.5629 | 690.5638 | 692.5647 | 4.0018 | 23 | 788.6083 | 790.6092 | 792.6102 | 4.0018 | 23 | 888.6538 | 890.6547 | 892.6556 | 4.0018 | 23 | 988.6993 | 990.7002 | 992.7011 | 4.0018
24 | 492.4737 | 494.4746 | 496.4756 | 4.0018 | 24 | 592.5192 | 594.5201 | 596.5211 | 4.0018 | 24 | 692.5647 | 694.5656 | 696.5665 | 4.0018 | 24 | 792.6102 | 794.6111 | 796.612 | 4.0018 | 24 | 892.6556 | 894.6566 | 896.6575 | 4.0018 | 24 | 992.7011 | 994.702 | 996.7029 | 4.0018
25 | 496.4756 | 498.4765 | 500.4774 | 4.0018 | 25 | 596.5211 | 598.522 | 600.5228 | 4.0018 | 25 | 696.5665 | 698.5674 | 700.5683 | 4.0018 | 25 | 796.612 | 798.6129 | 800.6138 | 4.0018 | 25 | 896.6575 | 898.6584 | 900.6593 | 4.0018 | 25 | 996.7029 | 998.7039 | 1000.7048 | 4.0018
26 | 398.431 | 400.4319 | 402.4328 | 4.0018 | 26 | 498.4765 | 500.4774 | 502.4783 | 4.0018 | 26 | 598.5219 | 600.5228 | 602.5237 | 4.0018 | 26 | 698.5674 | 700.5684 | 702.5692 | 4.0018 | 26 | 798.6129 | 800.6138 | 802.6147 | 4.0018 | 26 | 898.6584 | 900.6593 | 902.6602 | 4.0018
27 | 402.4328 | 404.4337 | 406.4346 | 4.0018 | 27 | 502.4783 | 504.4792 | 506.4801 | 4.0018 | 27 | 602.5237 | 604.5247 | 606.5256 | 4.0018 | 27 | 702.5692 | 704.5701 | 706.571 | 4.0018 | 27 | 802.6147 | 804.6156 | 806.6165 | 4.0018 | 27 | 902.6602 | 904.6611 | 906.662 | 4.0018
28 | 406.4346 | 408.4355 | 410.4365 | 4.0018 | 28 | 506.4801 | 508.481 | 510.4819 | 4.0018 | 28 | 606.5256 | 608.5265 | 610.5274 | 4.0018 | 28 | 706.571 | 708.572 | 710.5729 | 4.0018 | 28 | 806.6165 | 808.6174 | 810.6183 | 4.0018 | 28 | 906.662 | 908.6629 | 910.6638 | 4.0018
29 | 410.4365 | 412.4374 | 414.4383 | 4.0018 | 29 | 510.4819 | 512.4828 | 514.4838 | 4.0018 | 29 | 610.5274 | 612.5283 | 614.5292 | 4.0018 | 29 | 710.5729 | 712.5738 | 714.5747 | 4.0018 | 29 | 810.6183 | 812.6193 | 814.6202 | 4.0018 | 29 | 910.6638 | 912.6647 | 914.6657 | 4.0018
30 | 414.4383 | 416.4392 | 418.4401 | 4.0018 | 30 | 514.4838 | 516.4846 | 518.4855 | 4.0018 | 30 | 614.5292 | 616.5302 | 618.531 | 4.0018 | 30 | 714.5747 | 716.5756 | 718.5765 | 4.0018 | 30 | 814.6202 | 816.6211 | 818.622 | 4.0018 | 30 | 914.6657 | 916.6666 | 918.6675 | 4.0018
31 | 418.4401 | 420.441 | 422.4419 | 4.0018 | 31 | 518.4855 | 520.4864 | 522.4874 | 4.0018 | 31 | 618.531 | 620.5319 | 622.5328 | 4.0018 | 31 | 718.5765 | 720.5774 | 722.5783 | 4.0018 | 31 | 818.622 | 820.6229 | 822.6238 | 4.0018 | 31 | 918.6675 | 920.6683 | 922.6693 | 4.0018
32 | 422.4419 | 424.4428 | 426.4437 | 4.0018 | 32 | 522.4874 | 524.4883 | 526.4892 | 4.0018 | 32 | 622.5328 | 624.5338 | 626.5347 | 4.0018 | 32 | 722.5783 | 724.5792 | 726.5801 | 4.0018 | 32 | 822.6238 | 824.6247 | 826.6256 | 4.0018 | 32 | 922.6693 | 924.6702 | 926.6711 | 4.0018
33 | 426.4437 | 428.4446 | 430.4456 | 4.0018 | 33 | 526.4892 | 528.4901 | 530.491 | 4.0018 | 33 | 626.5347 | 628.5356 | 630.5365 | 4.0018 | 33 | 726.5801 | 728.5811 | 730.582 | 4.0018 | 33 | 826.6256 | 828.6265 | 830.6274 | 4.0018 | 33 | 926.6711 | 928.672 | 930.6729 | 4.0018
34 | 430.4456 | 432.4465 | 434.4474 | 4.0018 | 34 | 530.491 | 532.4919 | 534.4929 | 4.0018 | 34 | 630.5365 | 632.5374 | 634.5383 | 4.0018 | 34 | 730.582 | 732.5829 | 734.5838 | 4.0018 | 34 | 830.6274 | 832.6284 | 834.6293 | 4.0018 | 34 | 930.6729 | 932.6738 | 934.6747 | 4.0018
35 | 434.4474 | 436.4483 | 438.4492 | 4.0018 | 35 | 534.4929 | 536.4938 | 538.4946 | 4.0018 | 35 | 634.5383 | 636.5392 | 638.5401 | 4.0018 | 35 | 734.5838 | 736.5847 | 738.5856 | 4.0018 | 35 | 834.6293 | 836.6302 | 838.6311 | 4.0018 | 35 | 934.6747 | 936.6757 | 938.6766 | 4.0018
36 | 438.4492 | 440.4501 | 442.451 | 4.0018 | 36 | 538.4946 | 540.4955 | 542.4965 | 4.0018 | 36 | 638.5401 | 640.541 | 642.5419 | 4.0018 | 36 | 738.5856 | 740.5865 | 742.5874 | 4.0018 | 36 | 838.6311 | 840.632 | 842.6329 | 4.0018 | 36 | 938.6766 | 940.6775 | 942.6783 | 4.0018
37 | 442.451 | 444.4519 | 446.4528 | 4.0018 | 37 | 542.4965 | 544.4974 | 546.4983 | 4.0018 | 37 | 642.5419 | 644.5429 | 646.5438 | 4.0018 | 37 | 742.5874 | 744.5883 | 746.5892 | 4.0018 | 37 | 842.6329 | 844.6338 | 846.6347 | 4.0018 | 37 | 942.6783 | 944.6793 | 946.6802 | 4.0018
38 | 446.4528 | 448.4537 | 450.4547 | 4.0018 | 38 | 546.4983 | 548.4992 | 550.5001 | 4.0018 | 38 | 646.5438 | 648.5447 | 650.5456 | 4.0018 | 38 | 746.5892 | 748.5901 | 750.5911 | 4.0018 | 38 | 846.6347 | 848.6356 | 850.6365 | 4.0018 | 38 | 946.6802 | 948.6811 | 950.682 | 4.0018
39 | 450.4547 | 452.4556 | 454.4565 | 4.0018 | 39 | 550.5001 | 552.501 | 554.5019 | 4.0018 | 39 | 650.5456 | 652.5465 | 654.5474 | 4.0018 | 39 | 750.5911 | 752.592 | 754.5929 | 4.0018 | 39 | 850.6365 | 852.6374 | 854.6384 | 4.0018 | 39 | 950.682 | 952.6829 | 954.6838 | 4.0018
40 | 454.4565 | 456.4574 | 458.4583 | 4.0018 | 40 | 554.5019 | 556.5028 | 558.5037 | 4.0018 | 40 | 654.5474 | 656.5483 | 658.5493 | 4.0018 | 40 | 754.5929 | 756.5938 | 758.5947 | 4.0018 | 40 | 854.6384 | 856.6393 | 858.6402 | 4.0018 | 40 | 954.6838 | 956.6848 | 958.6857 | 4.0018
41 | 458.4583 | 460.4592 | 462.4601 | 4.0018 | 41 | 558.5037 | 560.5046 | 562.5055 | 4.0018 | 41 | 658.5493 | 660.5502 | 662.551 | 4.0018 | 41 | 758.5947 | 760.5956 | 762.5965 | 4.0018 | 41 | 858.6402 | 860.6411 | 862.642 | 4.0018 | 41 | 958.6857 | 960.6866 | 962.6875 | 4.0018
42 | 462.4601 | 464.461 | 466.4619 | 4.0018 | 42 | 562.5055 | 564.5065 | 566.5074 | 4.0018 | 42 | 662.551 | 664.5519 | 666.5529 | 4.0018 | 42 | 762.5965 | 764.5974 | 766.5983 | 4.0018 | 42 | 862.642 | 864.6429 | 866.6438 | 4.0018 | 42 | 962.6875 | 964.6884 | 966.6893 | 4.0018
43 | 466.4619 | 468.4628 | 470.4638 | 4.0018 | 43 | 566.5074 | 568.5083 | 570.5092 | 4.0018 | 43 | 666.5529 | 668.5538 | 670.5547 | 4.0018 | 43 | 766.5983 | 768.5992 | 770.6002 | 4.0018 | 43 | 866.6438 | 868.6447 | 870.6456 | 4.0018 | 43 | 966.6893 | 968.6902 | 970.6911 | 4.0018
44 | 470.4638 | 472.4647 | 474.4656 | 4.0018 | 44 | 570.5092 | 572.5101 | 574.511 | 4.0018 | 44 | 670.5547 | 672.5556 | 674.5565 | 4.0018 | 44 | 770.6002 | 772.6011 | 774.602 | 4.0018 | 44 | 870.6456 | 872.6466 | 874.6475 | 4.0018 | 44 | 970.6911 | 972.692 | 974.6929 | 4.0018
45 | 474.4656 | 476.4665 | 478.4674 | 4.0018 | 45 | 574.511 | 576.512 | 578.5128 | 4.0018 | 45 | 674.5565 | 676.5574 | 678.5584 | 4.0018 | 45 | 774.602 | 776.6029 | 778.6038 | 4.0018 | 45 | 874.6475 | 876.6484 | 878.6493 | 4.0018 | 45 | 974.6929 | 976.6938 | 978.6948 | 4.0018
46 | 478.4674 | 480.4683 | 482.4692 | 4.0018 | 46 | 578.5128 | 580.5137 | 582.5146 | 4.0018 | 46 | 678.5584 | 680.5592 | 682.5601 | 4.0018 | 46 | 778.6038 | 780.6047 | 782.6056 | 4.0018 | 46 | 878.6493 | 880.6502 | 882.6511 | 4.0018 | 46 | 978.6948 | 980.6957 | 982.6966 | 4.0018
47 | 482.4692 | 484.4701 | 486.471 | 4.0018 | 47 | 582.5146 | 584.5156 | 586.5165 | 4.0018 | 47 | 682.5601 | 684.561 | 686.562 | 4.0018 | 47 | 782.6056 | 784.6065 | 786.6074 | 4.0018 | 47 | 882.6511 | 884.652 | 886.6529 | 4.0018 | 47 | 982.6966 | 984.6975 | 986.6984 | 4.0018
48 | 486.471 | 488.4719 | 490.4728 | 4.0018 | 48 | 586.5165 | 588.5174 | 590.5183 | 4.0018 | 48 | 686.562 | 688.5629 | 690.5638 | 4.0018 | 48 | 786.6074 | 788.6083 | 790.6092 | 4.0018 | 48 | 886.6529 | 888.6538 | 890.6547 | 4.0018 | 48 | 986.6984 | 988.6993 | 990.7002 | 4.0018
49 | 490.4728 | 492.4737 | 494.4746 | 4.0018 | 49 | 590.5183 | 592.5192 | 594.5201 | 4.0018 | 49 | 690.5638 | 692.5647 | 694.5656 | 4.0018 | 49 | 790.6092 | 792.6102 | 794.6111 | 4.0018 | 49 | 890.6547 | 892.6556 | 894.6566 | 4.0018 | 49 | 990.7002 | 992.7011 | 994.702 | 4.0018
50 | 494.4746 | 496.4756 | 498.4765 | 4.0018 | 50 | 594.5201 | 596.521 | 598.5219 | 4.0018 | 50 | 694.5656 | 696.5665 | 698.5674 | 4.0018 | 50 | 794.6111 | 796.612 | 798.6129 | 4.0018 | 50 | 894.6566 | 896.6575 | 898.6584 | 4.0018 | 50 | 994.702 | 996.7029 | 998.7039 | 4.0018
51 | 498.4765 | 500.4774 | 502.4783 | 4.0018 | 51 | 598.5219 | 600.5228 | 602.5237 | 4.0018 | 51 | 698.5674 | 700.5684 | 702.5692 | 4.0018 | 51 | 798.6129 | 800.6138 | 802.6147 | 4.0018 | 51 | 898.6584 | 900.6593 | 902.6602 | 4.0018 | 51 | 998.7039 | 1000.7048 | 1002.7057 | 4.0018